5-Day SDI Lab Automation FDP

Participant Workbook

Exercises • Code Templates • Quick Reference • Day 5 AI Worksheets

Version 2.0

Mahaveer. Rajendra. Savanur  |  Head of Engineering & AI Systems

Spruha Build-in Solutions  |  Powered by Moku:Go — Liquid Instruments

Participant Name:  _______________________________

Institution:  _______________________________

Date:  _______________________________

Moku:Go IP Address:  _______________________________

Contents

Quick Start: Setup & config.py

Day 1: Oscilloscope & Waveform Generator

Day 2: AWG, Logic Analyzer & Multi-Instrument Mode

Day 3: Streaming, FFT & Signal Quality

Day 4: Industry Patterns & Capstone Projects

Day 5: MCC Session & AI Curriculum Design Worksheets

API Quick Reference Card (hardware-verified)

Troubleshooting Guide

1. Quick Start: Setup & config.py

Complete these steps before Day 1 (see the Pre-Workshop Checklist for the full list). Requires Python 3.10+.

pip install moku numpy matplotlib
mokucli instrument download   # one-time, ~500 MB
python -c "import moku; print(moku.__version__)"

Set your device IP — once, in one file

Every workshop script imports its connection settings from scripts/config.py. Find your device IP with "mokucli list" (or the device screen) and edit one line:

MOKU_IP = '192.168.0.101'   # <-- your Moku:Go IP

Connection test

from moku.instruments import Oscilloscope
from config import MOKU_IP, FORCE_CONNECT
 
osc = Oscilloscope(MOKU_IP, force_connect=FORCE_CONNECT)
try:
    print('Connected!')
finally:
    osc.relinquish_ownership()

CRITICAL:  Every script MUST call relinquish_ownership() in a finally block. Without it the Moku:Go stays locked and the next connection fails.

2. Day 1: Oscilloscope & Waveform Generator

Exercise 1.1 — Oscilloscope Data Logging  [30 min]

Objective: capture 10 frames, save each to CSV, compute Vpp statistics, generate a summary plot.

Setup: BNC loopback Output 1 → Input 1.  Script: day1/ex1_1_oscilloscope_logging.py

osc.set_frontend(channel=1, impedance='1MOhm',
                 coupling='DC', range='10Vpp')
osc.generate_waveform(1, 'Sine', amplitude=0.5,
                      frequency=1e3)
osc.set_timebase(-5e-3, 5e-3)
osc.set_trigger(type='Edge', source='Input1', level=0)
data = osc.get_data()   # dict of Python LISTS
vpp = max(data['ch1']) - min(data['ch1'])

REMEMBER:  get_data() returns Python lists — min()/max() work directly; convert with np.array() before numpy math.

My observations (Vpp across 10 frames)

Mean Vpp: ____________   Std dev: ____________   Stable? ____________

Exercise 1.2 — Waveform Generator Parametric Sweep  [30 min]

Objective: automate a 21-combination frequency × waveform sweep and measure the efficiency gain vs manual operation.

Script: day1/ex1_2_waveform_generator_sweep.py

API NOTE:  Square needs duty=, Ramp needs symmetry=, Pulse needs duty= — the API validates these per type.

My efficiency result

Automated time: __________ s    Manual estimate: __________ s    Gain: __________ × faster

3. Day 2: AWG, Logic Analyzer & MiM

Exercise 2.1 — AWG Custom Waveform Library  [30 min]

Objective: define 5 waveforms with numpy, save a JSON library, output each on the AWG.  Script: day2/ex2_1_awg_custom_waveforms.py

awg.generate_waveform(
    channel=1,
    sample_rate='Auto',
    lut_data=list(wf),   # Python LIST in [-1, 1]
    frequency=1e3, amplitude=1.0)

#1 MISTAKE:  Passing a numpy array as lut_data. Always wrap with list(), and normalize to [-1, 1] first.

Exercise 2.2 — Logic Analyzer Digital Patterns  [30 min]

Objective: generate a 4-pin SPI-like bus and capture it back. Script: day2/ex2_2_logic_analyzer_spi.py

HARDWARE-VERIFIED API (moku v4.1.2):  all pins go in ONE set_pattern_generator(channel=1, patterns=[...]) call, and pin mode is set_pin_mode(pin, 'PG1') — not 'Pattern'.

la.set_pattern_generator(
    channel=1,
    patterns=[{"pin": 1, "pattern": clock},
              {"pin": 4, "pattern": mosi}],
    divider=5)
la.set_pin_mode(1, 'PG1')

Exercise 2.3 — Multi-Instrument Mode Frequency Sweep  [45 min]

Objective: run WG (Slot 1) + Oscilloscope (Slot 2) simultaneously, sweep 10 Hz–100 kHz, plot the Bode response. Script: day2/ex2_3_mim_frequency_sweep.py

CRITICAL FORMAT:  connections is a LIST of DICTS. A plain dict silently loses duplicate keys and fails.

m = MultiInstrument(MOKU_IP, platform_id=2,
                    force_connect=True)
wg  = m.set_instrument(1, WaveformGenerator)
osc = m.set_instrument(2, Oscilloscope)
m.set_connections(connections=[
    dict(source='Slot1OutA', destination='Output1'),
    dict(source='Input1',    destination='Slot2InA'),
])
# In MiM 'Input1' is NOT a valid trigger source —
# use auto-trigger (just don't call set_trigger)
...
m.relinquish_ownership()   # on m — NEVER on wg/osc

My -3 dB cutoff (with RC filter inserted)

Measured fc: ____________ Hz    Theory 1/(2πRC): ____________ Hz

4. Day 3: Streaming, FFT & Signal Quality

Exercise 3.1 — Datalogger Streaming  [30 min]

Objective: stream 10 s at 10 kSa/s with live chunk statistics, save all samples to CSV.  Script: day3/ex3_1_datalogger_streaming.py

dl.start_streaming(duration=10, sample_rate=10e3)
while True:
    try:
        data = dl.get_stream_data()
        if data: process(data['ch1'])
    except Exception as e:
        if 'End of stream' in str(e):
            break        # NORMAL completion
        raise

Exercise 3.2 — FFT, SNR & THD  [45 min]

Objective: compute the FFT of a captured frame, find the dominant frequency, calculate SNR and THD, grade the signal. Script: day3/ex3_2_fft_snr_thd.py

BENCHMARKS:  SNR > 40 dB = clean signal. THD < 1% = pure tone. Loopback should beat both easily.

My signal quality result

Dominant freq: __________ Hz    SNR: __________ dB    THD: __________ %

5. Day 4: Industry Patterns & Capstone Projects

The 5 Automation Patterns

| Pattern | Description | Examples |
| --- | --- | --- |
| Parameter Sweep | Vary one input, measure response | Bode plot, V-I curves |
| Pass/Fail Testing | Compare against spec limits | QC, go/no-go |
| Continuous Monitoring | Stream & alert on thresholds | Environmental, burn-in |
| Calibration Routine | Known input → correction factor | Sensor cal, zeroing |
| Batch Testing | Same sequence, N devices | Production line, sorting |

Pattern Matching Exercise

Write the pattern(s) that fit each scenario, and why:

| Scenario | Pattern(s) — your answer |
| --- | --- |
| Sort 200 resistors into 1% / 5% / 10% bins |  |
| Track power supply drift over 48 hours |  |
| Measure amplifier gain 10 Hz – 100 kHz |  |
| Correct temp sensor offset via ice bath |  |
| Verify 50 PCBs meet voltage-regulation spec |  |

Final Capstone — choose one

| Option | System | Goal |
| --- | --- | --- |
| A | Component Characterization (MiM sweep → datasheet) | manual ~40 min → < 5 min |
| B | Signal Quality Monitor (streaming + alerts) | impossible manually → unattended |
| C | Production Tester (FFT pass/fail + certificates) | manual 8 min/unit → < 90 s |

All three exist as complete scripts in scripts/day4/ — run them, understand them, then extend them. Rubric: Automation 50 • Technical 30 • Documentation 20 (+10 bonus).

6. Day 5: MCC Session & AI Curriculum Design

Part A — Moku Custom Instrument (MCC) Session  [8:30–10:30]

You will complete a minimal custom FPGA instrument (threshold detector) inside the fixed CustomWrapper module. Files: scripts/day5/mcc_threshold_detector_skeleton.v + mcc_threshold_detector_host.py + mcc_experiment_guide.md.

Worksheet: the 5 differences from your FPGA lab

| Your FPGA lab | MCC — fill in what changes |
| --- | --- |
| Name your own top module |  |
| Choose your clock frequency |  |
| Define your own I/O ports |  |
| Raw integer signal values |  |
| JTAG / hardware programmer |  |

My threshold detector results

Threshold (Control0): __________    Crossings/sec (Status0): __________    Expected: __________

Part B — AI-Assisted Curriculum Mapping  [10:45–12:30]

Open NEW chats in BOTH Claude (claude.ai) and ChatGPT — free tiers are enough.

Paste scripts/day5/sdi_context_primer_prompt.txt into both. Wait for "Context loaded".

Pick ONE subject you teach and list its syllabus units below.

Fill the brackets in curriculum_mapping_prompt.txt and paste it.

Iterate: challenge weak mappings, fix invented topics. You are the judge.

My subject

Subject: ______________________  Course code: __________  Semester: __________

Key units: ___________________________________________________________________

Mapping quality check (tick before accepting)

☐ Every mapped experiment ties to a REAL unit of my syllabus

☐ Weak fits are labeled honestly (not forced)

☐ Bloom's level has a one-line justification I agree with

☐ Lab slots fit my actual term calendar

Part C — AI-Assisted Experiment Design  [1:30–4:30]

GPT: paste gpt_ideation_prompt.txt (filled in) → 3 ideas → pick ONE.

Claude: paste claude_design_prompt.txt with your chosen idea → full experiment document.

Iterate with Claude until it fits your syllabus and lab format.

VERIFY: cross-check every API call against Days 1–4. Anything unfamiliar gets "# TODO: verify against moku docs".

Trainer review → finalize → 2–3 volunteers present at 4:30.

Code verification checklist (tick every line)

☐ try/finally with relinquish_ownership() present

☐ get_data() treated as Python lists (np.array before numpy math)

☐ MiM: connections as list of dicts; cleanup on m; ≤ 2 instruments

☐ No invented method names — everything matches Days 1–4 patterns

☐ Signals within 30 MHz bandwidth; feasible in a 2–3 h lab slot

My deliverables

☐ Curriculum map (1–2 pages, HoD-ready)        ☐ One complete SDI experiment document

7. API Quick Reference Card (hardware-verified)

Oscilloscope

| Call | Notes |
| --- | --- |
| set_frontend(ch, impedance, coupling, range) | '1MOhm'/'50Ohm', 'AC'/'DC', '10Vpp'… |
| generate_waveform(ch, type, amplitude=, frequency=) | built-in WG; 'Sine', 'Square'(duty=), 'Ramp'(symmetry=)… |
| set_timebase(t1, t2) | seconds relative to trigger |
| set_trigger(type='Edge', source='Input1', level=0) | standalone only — NOT valid in MiM (auto-trigger works) |
| get_data() | dict of Python LISTS: 'time', 'ch1', 'ch2' |
| relinquish_ownership() | REQUIRED, in finally |

Arbitrary Waveform Generator

| Call | Notes |
| --- | --- |
| generate_waveform(ch, sample_rate='Auto', lut_data=, frequency=, amplitude=) | lut_data = Python list normalized to [-1, 1] |

Logic Analyzer (verified moku v4.1.2)

| Call | Notes |
| --- | --- |
| set_pattern_generator(channel=1, patterns=[…], divider=) | ALL pins in ONE call; patterns = [{"pin": n, "pattern": […]}] |
| set_pin_mode(pin, 'PG1') | 'PG1'/'PG2'/'I'/'X' — NOT 'Pattern' |
| get_data(include_pins=[1,2,3,4]) | returns 'pin1', 'pin2', … |

Datalogger

| Call | Notes |
| --- | --- |
| start_streaming(duration=, sample_rate=) | fixed-length stream |
| get_stream_data() | latest chunk; raises 'End of stream' at the end — that is NORMAL completion |
| enable_input(ch, enable=) | stream only the channels you need |

MultiInstrument (MiM)

| Call | Notes |
| --- | --- |
| MultiInstrument(ip, platform_id=2, force_connect=True) | platform_id=2 on Moku:Go (2 slots) |
| m.set_instrument(slot, Class) | returns the instrument object |
| m.set_connections(connections=[dict(source=, destination=)…]) | LIST of DICTS — never a plain dict |
| m.set_frontend(ch, …) | frontend via m in MiM, not instruments |
| m.relinquish_ownership() | releases ALL slots — only on m |

MiM routing nodes (Moku:Go)

| Node | Description |
| --- | --- |
| Input1 / Input2 | physical ADC inputs |
| Output1 / Output2 | physical DAC outputs |
| Slot1InA/B, Slot2InA/B | inputs to each slot |
| Slot1OutA/B, Slot2OutA/B | outputs from each slot |

8. Troubleshooting Guide

| Symptom | Solution |
| --- | --- |
| Connection refused / timeout | Check IP in config.py; same network; force_connect=True |
| Instrument busy / locked | Another session holds it — force_connect=True or wait ~60 s |
| ModuleNotFoundError: moku | pip install moku |
| Bitstream not found | mokucli instrument download (one-time, needs internet) |
| get_data() returns noise | Check BNC loopback and set_frontend range |
| MiM set_connections error | Must be a LIST of dicts, not a plain dict |
| MiM cannot set frontend | Use m.set_frontend() on the MultiInstrument |
| 'End of stream' exception | Normal streaming completion — catch it and exit the loop |
| Values clipped at rail | Signal exceeds range — increase range in set_frontend |
| AI invents API methods (Day 5) | Re-paste the SDI context primer in a NEW chat; verify all code |

Docs: https://apis.liquidinstruments.com/  •  Forum: https://forum.liquidinstruments.com/  •  Support: info@spruha.in