## Slide 1
DAY 1 OF 5
Foundations — Part 1
"From Manual to Automated — Getting Started"
5-Day SDI Lab Automation FDP • v2.0
Mahaveer. Rajendra. Savanur
Head of Engineering & AI Systems
Spruha Build-in Solutions | Powered by Moku:Go — Liquid Instruments

## Slide 2
Day 1 Agenda
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
1
8:00 – 9:00
Introduction & Workshop Overview objectives, teams, the case for automation
2
9:15 – 10:30
Environment Setup & Verification Python, moku, mokucli, connection test
3
10:45 – 12:30
Oscilloscope & Waveform Generator Automation live demos + Exercises 1.1 and 1.2
4
1:30 – 3:30
Practice, Bonus Challenges & Office Hours
5
3:45 – 5:00
Code Review, Q&A and Day 1 Wrap-Up

## Slide 3
Why Automate Test & Measurement?
Case study: a 100-point frequency sweep
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Manual Testing
~2 hours per sweep
Error-prone and inconsistent
Limited test coverage
No regression testing
Results live in a notebook
Automated Testing
~2 minutes per sweep
Perfectly repeatable
Full coverage, every run
Re-run any time, 24/7 unattended
CSV + plots, every run
The skill you are building this week: teaching Python to run your lab instruments for you.

## Slide 4
The ROI of Automation
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
40 hrs
one-time development investment
200 hrs
saved per year (100 runs × 2 hrs)
~2.5 mo
payback period
60×
faster per test run
Industry Applications
✓
Academic Research: IIT Madras lock-in amplifier experiments, 48-hour unattended datasets
✓
Manufacturing: automated product testing & QA with zero data-entry errors
✓
R&D: design validation and characterization
✓
Education: consistent, repeatable lab experiences for every student

## Slide 5
Moku:Go — Your Lab in a Box
Software-defined instrumentation: 15+ instruments in one device
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Hardware
2× Analog Inputs (BNC, 30 MHz, 125 MSa/s)
2× Analog Outputs (BNC, 20 MHz)
16× Digital I/O pins
Programmable supplies (some units)
USB-C / Ethernet / Wi-Fi
Software
Oscilloscope, WG, AWG, Logic Analyzer,
Datalogger… — defined in software
Multi-Instrument Mode: 2 at once
Python / MATLAB / LabVIEW APIs
Over-the-air updates
Why It Matters
One device replaces racks of equipment
Software-defined = future-proof
Same automation skills transfer to
any instrument with an API
Focus of this FDP: test & measurement AUTOMATION — not product training.

## Slide 6
Environment Setup & Verification
Everyone connected before lunch — this is the #1 Day-1 blocker
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Verify Python
python --version
needs 3.10+
1
Install moku
pip install moku
numpy matplotlib
2
Bitstreams
mokucli instrument
download (~500 MB)
3
Find your IP
mokucli list
or device screen
4
Test connect
set config.py IP
run connection test
5
One IP to rule them all: every script imports MOKU_IP from scripts/config.py — edit it once.
✓
Connection refused? Same Wi-Fi network as the Moku:Go, IP correct in config.py
✓
Device busy? force_connect=True (already set in config) or wait ~60 s
✓
Stuck? Raise your hand — we fix connections before moving on

## Slide 7
The Core Automation Pattern
Every script this week follows the same four steps
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
CONNECT
Establish connection
to the Moku:Go
1
CONFIGURE
Set instrument
parameters
2
ACQUIRE
Capture data
& measurements
3
CLEANUP
Release ownership
& save results
4
# The pattern every script follows:
osc = Oscilloscope(MOKU_IP, force_connect=True) # CONNECT
osc.set_timebase(-0.001, 0.001) # CONFIGURE
data = osc.get_data() # ACQUIRE
osc.relinquish_ownership() # CLEANUP
Skip CLEANUP and the device stays locked. try / finally is not optional.

## Slide 8
Oscilloscope Automation — Live Demo
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
day1/ex1_1_oscilloscope_logging.py
from moku.instruments import Oscilloscope
from config import MOKU_IP, FORCE_CONNECT
osc = Oscilloscope(MOKU_IP, force_connect=FORCE_CONNECT)
try:
osc.set_frontend(channel=1, impedance='1MOhm',
coupling='DC', range='10Vpp')
osc.generate_waveform(1, 'Sine',
amplitude=0.5, frequency=1e3)
osc.set_timebase(-5e-3, 5e-3) # +/- 5 ms window
osc.set_trigger(type='Edge',
source='Input1', level=0)
data = osc.get_data() # Python LISTS, not numpy
ch1 = data['ch1']
vpp = max(ch1) - min(ch1) # list ops work directly
print(f'Vpp = {vpp:.4f} V')
finally:
osc.relinquish_ownership()
Loopback: Output 1 -> Input 1 (BNC cable)
Key Concepts
API mirrors the GUI
every knob in the app is a
method call in Python
get_data() returns
dict of LISTS: 'time','ch1','ch2'
np.array() before numpy math
set_timebase()
seconds relative to trigger
Always call
relinquish_ownership()
in a finally block

## Slide 9
Exercise 1.1: Oscilloscope Data Logging
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Objective: Capture 10 oscilloscope frames, save each to CSV, compute statistics, create a summary plot
Requirements
✓
Configure timebase: ±5 ms
✓
Configure trigger: Edge, Input1, 0 V
✓
Capture 10 frames (0.5 s delay each)
✓
Save each frame to its own CSV file
✓
Print stats: Min, Max, Vpp per frame
✓
Create a summary plot of all 10 frames
Expected Output
Frame 1 | 0.9921 | -0.4961 | 0.4960
frame_01.csv saved
Frame 2 | 0.9842 | -0.4921 | 0.4921
frame_02.csv saved
...
Mean Vpp: 0.9881 V Std: 0.0040 V
Plot saved: oscilloscope_summary.png
Duration: 30 minutes | Teams of 2
Bonus Challenge: Add SPC limits — mean and ±3 sigma bounds on the summary plot
Script: day1/ex1_1_oscilloscope_logging.py • compare with expected_outputs/ex1_1_expected.png

## Slide 10
Waveform Generator Automation — Live Demo
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
day1/ex1_2_waveform_generator_sweep.py
from moku.instruments import WaveformGenerator
from config import MOKU_IP, FORCE_CONNECT
wg = WaveformGenerator(MOKU_IP, force_connect=FORCE_CONNECT)
try:
# one waveform, three lines
wg.generate_waveform(channel=1, type='Sine',
amplitude=1.0, frequency=1000)
# automation = the same call in a loop
for freq in [100, 1000, 10000, 100000]:
wg.generate_waveform(channel=1, type='Sine',
amplitude=1.0, frequency=freq)
time.sleep(0.5)
# Square needs duty=, Ramp needs symmetry=
wg.generate_waveform(channel=1, type='Square',
amplitude=2.0, frequency=5000,
duty=30)
finally:
wg.relinquish_ownership()
The stimulus side — combined with the scope: stimulus-response testing
Key Concepts
Waveform types
'Sine', 'Square', 'Ramp',
'Triangle', 'Pulse', 'DC'
Per-type params
Square / Pulse need duty=
Ramp needs symmetry=
The whole idea
21 manual steps become
a 3-line loop

## Slide 11
Exercise 1.2: Automated Waveform Test Sequence
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Objective: Automate a frequency × waveform sweep and measure your efficiency gain vs manual operation
Requirements
✓
Test matrix: 7 frequencies (100 Hz – 100 kHz)
✓
Waveforms: Sine, Square, Ramp = 21 tests
✓
Log every test to wg_sweep_log.csv
✓
Time the run; compare vs ~30 s per manual change
✓
Print the efficiency report
Expected Output
Test 21/21: Ramp @ 100000 Hz
EFFICIENCY REPORT
Total tests: 21
Automated time: 12.4 s
Manual estimate: 630 s (10.5 min)
Efficiency gain: 50.8x faster
Duration: 30 minutes | share your × number
Bonus Challenge: Add amplitude sweep 0.1 – 2.0 V in 0.1 V steps; recompute total tests and time saved
Script: day1/ex1_2_waveform_generator_sweep.py

## Slide 12
Code Review & Best Practices
Afternoon session — bring your Exercise 1.1 / 1.2 code
Day 1: Foundations — Part 1
5-Day SDI Lab Automation FDP • v2.0
Powered by Moku:Go — Liquid Instruments
Organization
Imports at top
Config constants next
Functions for reuse
Main logic at the bottom
Error Handling
try / except / finally
Cleanup ALWAYS in finally
Print errors, don't swallow them
Data Management
CSV with header rows
Consistent file naming
Timestamps in logs
Plot Quality
Axis labels with units
Titles and grids
dpi=150 for reports
Efficiency
Minimal sleep times
Save incrementally on long runs
Settle ≥ 50 ms after waveform change
Reuse
config.py for shared constants
One pattern, many experiments
Your scripts become templates

## Slide 13
Day 1 Complete!
What You Achieved Today
✓
Set up the Python + Moku:Go development environment
✓
Learned the Connect → Configure → Acquire → Cleanup pattern
✓
Automated Oscilloscope capture with CSV logging & statistics
✓
Automated Waveform Generator parametric sweeps
✓
Measured your own efficiency gain vs manual operation
Tomorrow — Day 2 Preview
▸ Arbitrary Waveform Generator
▸ Logic Analyzer & digital patterns
▸ Capstone: Frequency Response Tester
▸ Multi-Instrument Mode introduction
▸ Internal routing & connections
Homework
1. Review today's code — ensure both exercises run cleanly
2. Read the workbook pages on AWG and Logic Analyzer
3. Prepare questions for tomorrow's sessions
Spruha Build-in Solutions | Powered by Moku:Go — Liquid Instruments